Method: Determination of chlorophyll a concentrations

1-Measure directly with the small blue Turner fluorometer (Calibration is done!)

- Calibration of the fluorometer

- Switch on at least 20 min. before use

- Calculate the concentration of chlorophyll a in the ethanol extract that was measured using a spectrofluorometer, as follows:

Cwater=Cextract× vV

Cwater =concentration of chlorophyll a in water (µg l-1)

Cextract=concentration of chlorophyll a in the extract (µg l-1) (measured in fluorometer)

v = volume of ethanol (l)

V = filtered volume of water (l)

Measure a blank with only ethanol (ask your instructor, or read the lazy-dog).

Place the extract in the fluorometer and set the standard value to the extract concentration. If the extract is too concentrate (value > 350), dilute the extract and start again. The best is to calibrate the fluorometer with an extract of similar concentration range as your samples.

2- Fluorometric determination of chlorophyll a

- Filter onto glassfiber filters (make sure that you are in a low light environment, and that you note the EXACT volume you filter- use a graduated pipette)

- Filter carefully a known volume (10-50 ml) with low vacuum (< 0.3 bar)

- Let the filter dry

- Insert filter in labeled glass tubes (pencil and paper tape…)

- Add 4 ml Ethanol

- Add caps and store in dark (room temp) for extraction (6-8h)

- Fluorometer ON (20-30 min before measurements)

- Use the appropriate calibration program (see with your instructor)

- Use the round vial and appropriate holder (ask instructor)

- Measure a blank sample with only ethanol (ask your instructor, or read the lazy-dog).

- Transfer extract into the 2 ml glass vial

- Dry carefully with paper

- Measure the fluorescence/ chlorophyll a concentration in the extract (important to remember while calculating the chla concentrations in the cultures)

Volume extract = 4 ml

Volume filtered = X ml

Concentration in extract = known chla µg l-1

Calculate concentration in culture.